## Slide 1: The ERP Boot Camp
Setting up and Running an ERP Lab

## Slide 2: Recording Chamber
Do you really need one?
Probably not if:
  You’re looking at slow components and can low-pass filter with a 50% cutoff at 30 Hz
  And you’re not near any major source of electrical noise
    Elevators, centrifuges, power transformers, ventilation fans
  You don’t care about gamma oscillations
  Or you have active electrodes
They’re good for keeping subjects focused
They tend to get warm, so it may actually be better not to have one if skin potentials are a major source of noise
You can build one from 2x4’s and copper screen
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 3
Courtesy of Lynne Reder
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 4: Seating
Key points:
  Comfortable to avoid muscle noise
  Don’t want subjects to fall asleep
  Don’t want electrodes to rest on anything
Recliners were once common
  Not good if you have electrodes over the back of the head
I recommend high-quality office chair
  Glides rather than wheels
  Mark the floor
I haven’t had much luck with chin rests
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 5: Response Devices
Need to be held in a comfortable position
  Don’t want subject holding arms up
  Standard computer keyboards are bad
Game controllers work well
  Mass-produced -> reliable
Constant and variable timing errors are possible
  RT is so variable that a bit of timing variability will usually have virtually no impact (unless you are looking at response-locked averages)
    Ulrich, R., & Giray, M. (1989). Time resolution of clocks: Effects on reaction time measurement—Good news for bad clocks. British Journal of Mathematical and Statistical Psychology, 42, 1-12.
  EMG for best response timing
  Can measure timing errors by putting a mic next to device and recording “click” along with event code
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 6: Hints for Running Subjects
~60 minutes of “run time” per session
  More for interesting experiments
  Whole session is about 3 hours
Runs of 4-6 minutes with 2-3 20-second breaks
  Less makes it inefficient to deal with electrodes, etc.
  More leads to fatigue
  Some labs do all-day sessions with lot of breaks
Dilution Rule: Don’t dilute good data with bad data
  Adding noisy trials doesn’t improve the S/N ratio
Watch the EEG throughout the session
  Look for artifacts, bad connections, etc.
Watch the subject with a video camera
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.
Backpack for experimenter
Cell phone for subject or experimenter
Sound proofing versus masking noise

## Slide 7: Hints for Running Subjects
Happiness Rule: A happy subject is a good subject
  Compliance with task
  Compliance with artifact control instructions
  Less noise
Talking to subjects
  Treat subject like a person, not like a piece of meat
  Chat while putting on electrodes (or video)
    Tell them exactly what will happen -- this reduces stress
  Chat during breaks
  Note: Some subjects don’t want to talk -- that’s OK
Keeping subjects happy
  Food and drink (is caffeine a confound?)
  Eye drops (single-use)
  Music
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 8: Looking at the Data
Do a fairly complete analysis of the first subject’s data before running anyone else
  All the main comparisons among ERP waveforms
  Accuracy (and RT if recorded) for each main condition
  There may be a serious problem with event codes, etc.
  Nothing focuses the mind quite as much as real data
Take a look at the individual subjects and the grand averages every 3-4 subjects
  Grand averages will give you more power to see something funky in the data
  But don’t get too freaked out if the results look a little funny or aren’t conforming to your predictions
  Be especially concerned about “impossible” results (e.g., effects that consistently begin before time zero)
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 9: Looking at the Data
When you do stats
  Have your stats program print a table of means
  Compare the means with your grand averages
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 10: Ethical Issues
Everything that applies to behavioral experiments plus…
  Risk of disease transmission
    High impedance helps (because abrasion is not needed)
    Need thoughtful disinfection (even for high impedance)
  Risk of electrical shock
    Optical isolation and/or battery power
  Headache from electrode cap
  Gel in hair
  Long duration of experiment
  Claustrophobia
  Concerns about privacy of EEG data
Providing clear information in advance is the best way to prevent problems
Also: Don’t fish for p values
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 11: CRT Basics
Notes: Figure 8.3 from Luck, Steven J. (2005). An introduction to the Event-Related Potential Technique. Cambridge, MA: MIT Press.© MIT Press. This material may be used for nonprofit research and education purposes only, and it may not be reprinted or distributed in any form including print and electronic forms.
Other content: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 12: Stimulus Presentation
Testing the timing of event codes
  Digitize at ~1000 Hz
  Point some kind of light-sensitive device to the video monitor and connect to digitization system
    Measure from the upper left corner of the monitor
    Also try other locations to assess delay
  Present stimuli with event codes and record the stimulus
  See when the stimuli are actually presented relative to the event codes
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 13: CRT Test
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 14: LCD Test
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 15: Auditory Stimulus Presentation
Auditory artifacts
  Speaker in headphones may induce a current in your electrodes/wires
  Post-auricular muscle twitch
Testing the timing of event codes
  Digitize at fastest possible rate (e.g., 10KHz)
  Connect auditory output (or a microphone) to the digitization system
    You might want to use a square-wave tone or a 50-Hz sine wave
  Present stimuli with event codes and record the stimulus
  See when the stimuli are actually presented relative to the event codes
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.

## Slide 16: Auditory Stimulus Presentation
Look for constant and variable delays
Notes: © S. J. Luck. All Rights Reserved. May be used for nonprofit educational purposes if this copyright notice is included. Permission must be obtained from the copyright holder(s) for any other use.